{"text_format":"1","workbook":"xlsx","n_sheets":1,"sheets":["VALEPORT SVX2"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"VALEPORT SVX2","index":0,"n_rows":114,"n_cols":12,"rows_kept":114,"cut":false,"hidden":false,"merged":[[51,1,54,1],[51,2,54,11],[47,1,47,9],[47,10,47,11],[48,1,48,9],[48,10,48,11],[49,1,49,9],[49,10,49,11],[44,1,44,9],[44,10,44,11],[45,1,45,9],[45,10,45,11],[46,1,46,9],[46,10,46,11],[41,1,41,9],[41,10,41,11],[42,1,42,9],[42,10,42,11],[43,1,43,9],[43,10,43,11],[38,1,38,9],[38,10,38,11],[39,1,39,9],[39,10,39,11],[40,1,40,9],[40,10,40,11],[34,1,34,8],[36,1,36,8],[36,10,36,11],[37,1,37,9],[37,10,37,11],[29,1,29,8],[30,1,30,8],[31,1,31,8],[32,1,32,8],[33,1,33,8],[23,1,23,8],[24,1,24,8],[26,1,26,8],[27,1,27,8],[28,1,28,8],[25,1,25,8],[18,1,18,8],[19,1,19,8],[20,1,20,8],[21,1,21,8],[22,1,22,8],[12,1,12,8],[13,1,13,8],[14,1,14,8],[15,1,15,8],[17,1,17,8],[16,1,16,8],[3,6,3,10],[4,6,4,10],[6,1,6,2],[6,3,6,4],[6,6,6,7],[6,8,6,9],[9,6,9,7],[9,8,9,9],[7,1,7,2],[7,3,7,4],[7,6,7,7],[7,8,7,9],[8,1,8,2],[8,3,8,4],[8,6,8,7],[8,8,8,9]]}
[]
[]
[]
[null,null,null,null,null,null,"VALEPORT SVX2"]
[null,null,null,null,null,null,"INSPECTION  CHECKLIST"]
[]
[null,"Customer: ",null,null,null,null,"Date of Inspection: ",null,{"$f":"NOW()","v":"2021-07-21T11:43:57.860"}]
[null,"Job Number: ",null,null,null,null,"Pre/Post Rental: "]
[null,"Asset Number: ",null,null,null,null,"Serial Number: "]
[null,null,null,null,null,null,"Checked by: ",null,{"$f":"M11Username()","v":{"$error":"#NAME?"}}]
[]
[null,null,null,null,null,null,null,null,null,"PASS","FAIL","N/A"]
[null,"Check unit for damage, particularly attention the conductivity cell, temperature sensor & SV sensors (for cracks, impact or sensors covered in dirt)."]
[null,"Check connector and locking collar for corrosion/damage."]
[null,"Remove pressure sensor plastic protective cap & ensure pressure sensor is free from debris /  damage to sensor diaphram"]
[null,"Check the Mounting frame has no corrosion or pitting on its metal frame + the unit is secured tightly and positioned correctly in the frame."]
[null,"Record calibration date of unit:"]
[null,"Ensure calibration information is up to date in AX and the latest calibration certificate is stored in the Assets folder on the Unique network"]
[null,"Check software version is current as per Valeport’s website & record:"]
[null,"Check manual revisions are current as per Valeport’s website & record:"]
[null,"Record baud rate & ensure unit is marked accordingly."]
[null,"With the unit running on batteries only run the Datalog Express or Datalog X2 software & establish comms with the unit. Preferably use Datalog Express as data can be graphed real time in the software."]
[null,"Check Internal Batteries have enough voltage for unit to operate (>9VDC)"]
[null,"Place the unit in the test tank & start the unit running in direct mode."]
[null,"Allow the readings to stabilise for 15 minutes & ensure that the measured sound velocity is the same as the calculated sound velocity to within +0.25m/s."]
[null,"Also deploy a freshly calibrated SVX2 If possible in the tank at the same position as the current one. Ensure the readings of both are within spec (give both systems at least 15 minutes to stablise in the tank before verifying readings)."]
[null,"Remove Unit from test tank."]
[null,"Check pressure sensor readings are correct when compared to atmospheric pressure (Tare off)."]
[null,"Turn Tare function on to zero the depth reading, place in test tank & ensure depth changes accordingly."]
[null,"Start the unit running in self recording mode and log data for 15 minutes."]
[null,"Ensure data can be downloaded correctly & erase the unit’s memory."]
[null,"Ensure unit functions when running on 24VDC external power (displayed battery voltage should increase accordingly in the software)."]
[null,"Record the battery voltage here (should be greater than 11.5V):"]
[null,"Record lithium battery voltage here (should be greater than 2.8V):"]
[null,"Ensure Unit is clean & presentable."]
[]
[null,null,null,null,null,null,null,null,null,null,"Tick as required"]
[null,"Package unit in its blue plastic case / yellow accessories bag complete with:"]
[null,"3M Y-Lead"]
[null,"10-Pin Switch Blanking Plug"]
[null,"Stainless Steel Strop c/w  3 x Shackles"]
[null,"Calibration Certificate"]
[null,"Spare O-Rings/Anti Extrusion Rings/Screws/Test Loop"]
[null,"USB Serial Convertor "]
[null,"USB Serial Adapter Drivers CD"]
[null,"USB Stick c/w software + Manuals. Record Software + Manual Versions:"]
[null,"12 or 24VDC DC Power Supply c/w banana plugs"]
[null,"UK, EU + US AC Power Lead For the DC Power Supply"]
[null,"Hard Copy Of Software + Hardware manuals. Record Versions:"]
[null,"Transit Case"]
[]
[null,"Comments:"]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[]
[null,"Pre","Post","Both"]