{"text_format":"1","workbook":"xlsx","n_sheets":1,"sheets":["COMPATT 6"],"not_rendered":{"charts":0,"images":1,"macros":1,"external_links":0,"defined_names":1,"comments":0,"pivot_tables":0},"unopened":[],"bounds":{"document_bytes":16777216,"sheet_cells":100000,"rows_first":2000,"rows_last":500,"rows_between":2000,"halvings":0,"cell_chars":16384,"cell_head_chars":4096},"cells_cut":0,"cells_raw":0,"read":"whole","date_system":"1900"}

{"sheet":"COMPATT 6","index":0,"n_rows":52,"n_cols":12,"rows_kept":52,"cut":false,"hidden":false,"merged":[[26,1,26,8],[34,1,34,11],[29,1,29,8],[30,1,30,8],[31,1,31,8],[32,1,32,8],[33,1,33,8],[19,1,19,8],[20,1,20,8],[21,1,21,8],[27,1,27,8],[28,1,28,8],[22,1,22,8],[23,1,23,8],[24,1,24,8],[25,1,25,8],[7,1,7,2],[7,3,7,4],[7,6,7,7],[7,8,7,9],[8,1,8,2],[8,3,8,4],[8,6,8,7],[8,8,8,9],[16,1,16,8],[17,1,17,8],[18,1,18,8],[9,6,9,7],[9,8,9,9],[12,1,12,8],[13,1,13,8],[14,1,14,8],[15,1,15,8],[3,6,3,10],[4,6,4,10],[6,1,6,2],[6,3,6,4],[6,6,6,7],[6,8,6,9],[35,2,35,9],[44,2,44,9],[44,10,44,11],[42,2,42,9],[42,10,42,11],[43,2,43,9],[36,2,36,9],[41,10,41,11],[36,10,36,11],[37,2,37,9],[37,10,37,11],[38,2,38,9],[38,10,38,11],[39,2,39,9],[35,10,35,11],[40,10,40,11],[45,10,45,11],[43,10,43,11],[39,10,39,11],[45,2,45,9],[51,1,54,1],[51,2,54,11],[46,10,46,11],[47,10,47,11],[48,10,48,11],[48,2,48,9],[46,2,46,9],[47,2,47,9],[49,10,49,11],[49,2,49,9],[41,2,41,9],[40,2,40,9]]}
[]
[]
[]
[null,null,null,null,null,null,"COMPATT 6"]
[null,null,null,null,null,null,"INSPECTION  CHECKLIST"]
[]
[null,"Customer: ",null,null,null,null,"Date of Inspection: ",null,{"$f":"NOW()","v":"2020-11-12T12:58:16.530"}]
[null,"Job Number: ",null,null,null,null,"Pre/Post Rental: "]
[null,"Asset Number: ",null,null,null,null,"Serial Number: "]
[null,null,null,null,null,null,"Checked by: ",null,{"$f":"Username()","v":"josh"}]
[]
[null,null,null,null,null,null,null,null,null,"PASS","FAIL","N/A"]
[null,"Record endcap asset number:"]
[null,"Record endcap type: "]
[null,"Ensure calibration is valid & record calibration date(s): \nPressure (Valid for 1 Year):                    \nSound Velocity (Valid For 2 Years):                               \nInclinometer (Valid For 2 Years): "]
[null,"Carefully inspect the unit for damage paying particular attention to the transducer."]
[null,"Check connectors for corrosion/damage."]
[null,"Check condition of anodes & replace if required."]
[null,"Check operation of the pressure relief valve."]
[null,"Check pressure inlet port is free from debris."]
[null,"Connect compatt to PC using serial cable & establish comms using the terminal program."]
[null,"Record battery voltage & percentage remaining (for pre-hire checks ensure battery voltage is sufficient for the duration of the job)."]
[null,"Check CPU & DAS Firmware are current as per software register & record(manual command FVI)\nCPU:\nDAS:"]
[null,"Check address matches label on outside of unit & record:"]
[null,"Pre hire only: If job is for multiple compatts/RovNavs ensure there are no conflicting addresses."]
[null,"Using another enabled compatt/iWand carry out an acoustic test & verify correct operation."]
[null,"If fitted check operation of the release (open & close)."]
[null,"Record sensors fitted & verify correct operation temp/pressure/Inc (MEMs)\nInt PRT (Temp) - \nRelease - \nInt RS485 Strain Pressure - \nDigiquartz Pressure -\nInt SPI Accel/Inclin -\nSound Velocity - \n"]
[null,"For units fitted with digiquartz sensors compare pressure reading with atmospheric & verify the unit is in specification."]
[null,"For SV sensors they must be checked alongside a freshly calibrated SV sensor (eg MiniSVS) in test tank. Allow 15 minutes of stabilising in exactly the same position as each other in test tank. The Endcap SV readings must be within 0.3M/S of the freshly calibrated SV. \nTake screenshot of results to prove SV readings and save in endcap asset folder."]
[null,"Produce local test report, check for errors & attach copy to this checklist."]
[null,"For units going into storage disconnect battery pack ensuring labelled as disconnected."]
[null,"Ensuring O-Rings are clean refit connector blanking cap & tighten using tool."]
[null,"Ensure unit is clean & presentable."]
[]
[null,"Qty","Package unit in transit case complete with:",null,null,null,null,null,null,null,"Tick as required"]
[null,1,"Sonardyne 6G Compatt"]
[null,1,"3000M Release End Cap"]
[null,1,"6G SV Endcap"]
[null,1,"3000M Digi/Inc Endcap"]
[null,1,"7000M Digi/Inc Endcap"]
[null,1,"Blanking Plug C/w O Ring"]
[null,1,"Blanking Plug Removal Tool"]
[null,1,"DAS/CPU Serial Y Lead"]
[null,1,"USB Stick c/w latest software, manuals and firmware"]
[null,1,"Calibration Certificate"]
[null,2,"Shackle"]
[null,1,"Test Report"]
[null,1,"6G Compatt Service Kit"]
[null,1,"Transit Case"]
[]
[null,"Comments:"]